## Slide 1: Cryostat Design Vittorio Parma CERN Technology Department, Magnets, Superconductors and Cryostats Group CERN Accelerator School «Superconductivity for Accelerators» Ettore Majorana Foundation and Centre for Scientific Culture Erice, Italy, 24 April – 4 May 2013

## Slide 2: Content
Introduction to cryostats
Cryostat requirements
Heat transfer for cryostats:
  Solid conduction
  Residual gas conduction
  Radiation, MLI protection, thermal shielding
Cryogenics
Heat intercepts
Insulation vacuum and construction issues
Mechanical considerations and construction codes
Supporting systems
Over-pressure safety issues
1st hour
2nd hour

## Slide 3: A bit of History
Cryostat (from cryo meaning cold and stat meaning stable): “a device used to maintain at cryogenic temperatures samples or devices mounted within the cryostat”
Dewar invents the “dewar”, 1892, London
A dewar: the first performing cryostat
  silvered, double-walled, glass vacuum vessel to contain cryogenic liquids
  J.Dewar: 1st liquefaction of H2 in 1897
  …but did not manage liquefaction of He, achieved by H.Kamerlingh Onnes in 1908
Glassblowers: the “enabling technology” of the epoque:
  J.Dewar did not patent his invention…
  H.K.Onnes created the “Leidse Instrumentmakersschool” (still existing!), and industrialized cryostats
Sir James Dewar (1842-1923)

## Slide 4: Dewars on “Google images”

## Slide 5: Applications
Cryostats
Liquid storage/transport
SC devices for accelerators
Test cryostats
Accelerator devices
SC devices for medical applications
Commercial vertical test cryostat
SC magnet cryostats
SCRF cavities cryostats
Specific vertical test cryostats

## Slide 6: A few examples at CERN
The High Field Magnet test cryostat
The HIE Isolde cryo-modules
The LHC cryostats

## Slide 7: Cryostats for SC devices for accelerators: A multidisciplinary activity
Cryostat
Design
Superconductivity
Heat Transfer
Mechanics
Cryogenics
Vacuum
System
Integration

## Slide 8: Cryostat requirements

## Slide 9: Functions
Two main functions:
Mechanical housing of cryogenic devices (supporting systems):
  Supporting of (sometimes heavy) SC devices
  Accurate & reproducible positioning (almost always)
  Precise alignment capabilities (SC devices in accelerators)
Thermal efficiency of the cryostat (heat loads as low as possible):
  Cooling capability (SC device, thermal shields and heat intercepts)
  Insulation vacuum (SC devices “hidden” in vessels)
  Thermal radiation shielding (screens, MLI)
  Low heat conduction (low thermal conductivity materials)
Often conflicting,  calls for trade off design solutions
Many other complementary functions…:
  Integration of cryogenic equipment (ph.separators, valves, etc.)
  Cryogenic cooling piping and interfaces to cryoplant
  Integration of Beam instrumentation (e.g.BPMs, BLMs,etc.)
  Instrumentation wires feed-throughs (control/diagnostics)
  magnetic shielding from/to environment (e.g. SCRF cavites, magnets)
  Maintainability (access ports)
  Handling and transport features
  …

## Slide 10: Mechanical Housing example of LHC
Low heat in-leaks support posts
Alignment jacks under a dipole
Alignment targets
LHC Main Cryostat
(Cross-Section)
Vacuum Vessel
Magnet Support Posts (GFRE)
External supports (jacks)

## Slide 11: Geometrical Stability: survey measurements
Mean: +0.1mm; St.dev.: 0.17mm
Mean: +0.08mm; St.dev.: 0.11mm
Survey measurements
with laser tracker
Transversal movements
Cryo-dipole
Vertical movements
Cryo-dipole
100 m descent to the tunnel
ROCLA vehicle
Tunnel transportation
Cold mass stability w.r.t. fiducials measurements on 20 cryo-dipoles
After transport to the tunnel
Quad CM positional stability and reproducibility at cold

## Slide 12: Thermal efficiency
Cooling of SC device:
Ensure operating T: cryo scheme, fluid distribution and heat transfer. Strongly coupled to cryoplant and cryo distribution system
Heat loads management:
External heat in-leaks
  Radiation
  Residual gas conduction
  Solid conduction
Internal heat sources:
  Joule heating (SC magnet splices)
  BCS residual resistance (RF cavities)
Beam-induced heat:
  Synchrotron radiation
  Beam image currents
  Photoelectrons (e-cloud)
  Mitigation measures:
  70 K shielding, MLI
  Vacuum < 10-4 Pa
  Low conductivity materials (non metallic), heat intercepts
  Special brazing (resistance < a few nW)
  Q enhancement, pulsed operation
  5-20 K beam screens
  Cu plated beam screens
  5-20 beam screens
static
dynamic

## Slide 13: Thermal efficiency example of LHC
Temperature | 50-75 K | 4.6-20 K | 1.9 K LHe | 4 K VLP
Static heat inleaks* | 7.7 | 0.23 | 0.21 | 0.11
Resistive heating | 0.02 | 0.005 | 0.10 | 0
Beam-induced nominal** | 0 | 1.58 | 0.09 | 0
Total nominal | 7.7 | 1.82 | 0.40 | 0.11
* no contingency |  |  |  | 
** Breakdown | nominal
Synchrotron radiation | 0.33
Image current | 0.36
Beam-gas Scattering | 0.05
Photoelectron | 0.89
LHC budgeted distributed steady-state heat loads [W/m]
HeII

## Slide 14
Heat Transfer for cryostats

## Slide 15: Heat transfer: General
Solid conduction:
Thermal radiation:
(with and without MLI)
Viscous gas conduction and natural convection: (Negligible with good insulation vacuum, < 10-4 Pa)
Gas conduction: molecular regime
Between cylinders:

## Slide 16
Thermal conduction

## Slide 17: Thermal Conduction
When a T gradient exists in a body, there is a heat transfer from the high T region to the low T region (Fourier Law):
For one-dimensional problems (ex. a bar or tube):
k is the thermal conductivity (W/mK-1), normally a function of P,T, material structure, non-homogeneity, anisotropy (ex. Composite materials).
k is strongly T-dependent and non-linear at low T
“good conductors” vs. “poor conductors”  k range ~ 5 orders of magnitude
Note: sometimes conductivity denoted by λ.
Tw
Tc

## Slide 18: Thermal conductivity in solids (& metals)
In metals, the electron contribution dominates.
The movement of conduction electrons is impeded by scatter: interactions with phonons, and interactions with impurities/imperfections. We can introduce thermal resistivities:
The conductivity is attributed to the movement of conduction electrons (“electron gas”), ke, and the effects of phonon lattice vibrations, kl.
Therefore for metals, the conductivity can be expressed as:
And has a maximum conductivity:
Therefore, for metals:
Kmax shifts at higher T with increasing impurity (see coppers and aluminiums )
The maximum vanishes for highly impure alloys (see steels) and in these cases impurity scattering dominates phonon scattering, thus at T < RT:
For metals, from electron conduction theory and analogy with electrical diffusion  Wiedemann-Franz law :
  Good agreement at T<< and T>> RT
  Better agreement from T<< to T>> with increasing impurities
  Electrical resistivity (ρe) easier to measure than thermal conductivity
e-
N+

## Slide 19: Thermal conductivity for Aluminium alloys

## Slide 20: Electrical resistivity for some metals

## Slide 21: The conduction equation (unidirectional case)
Longitudinal
conduction
Internal
heat
generation
Thermal
inertia
Or, if k ~ const.with T and introducing α thermal diffusivity:
(Fourier law of heat conduction)
(change of internal energy)
α allows evaluating the characteristic propagation time τ of a thermal

## Slide 22
No local heat deposition and steady-state
If k constant with T:
If k = k(T):
Thermal conductivity
Integral (conductance)
(Linear solution)
For impure metals (ex.steels) at low T:
(quadratic solution)
T
x
T0
TL
L

## Slide 23
Electrical network analogy
The inverse of the thermal conductance  thermal resistance:
In both cases we can recognize an analogy with the electrical resistance (replace q with I, T with V):
a) For constant k:
b) For variable k, define an average value kAV:
We can therefore model a complex thermal conductivity problem by elementary thermal resistances Ri, and solve the network by using Kirckhoff’s laws.

## Slide 24: Thermal conductivity data for selected materials

## Slide 25: Thermal conductivity of various materials
Spans over about 5 orders of magnitude

## Slide 26: Thermal conductivity integrals (conductance) for some materials [W/m]
Highest T (Lowest T =4.2 K) | 20 K | 80 K | 290 K
OFHC Copper | 11000 | 60600 | 152000
DHP Copper | 395 | 5890 | 46100
Aluminium 1100 | 2740 | 23300 | 72100
Aluminium 2024 | 160 | 2420 | 22900
Stainless steel AISI 304 | 16,3 | 349 | 3060
Typical Glass-fiber/Epoxy Composite G-10 | 2 | 18 | 153
Tw
4.2K
Qc

## Slide 27: …more conductivity integrals

## Slide 28: Thermal conductivity in composite materials
Generally non-conductors (little electron conduction), essentially phonon driven
Anisotropic structure, (fibers/matrix) with constituents-specific thermal conductivity properties
Generally homogeneous at macroscopic scale, but non-homogeneous at microscopic level (interface effects)
Conductivity highly depends on:
  Material (fiber) orientation
  Ratio between fibre and matrix (Vf)
For Glass-fiber/epoxy matrix composites (of wide interest in cryostat applications):
Glass is the “conductive” material and also the “structural” constituent
Epoxy is the “isolating” material and also the “less structural” constituent
Vf typically around 40-60%
Conductivity calculation difficult, opt for experimental measurements
Conductivity measurement of candidate GFRE materials for LHC supports

## Slide 29: Thermal conduction with uniform heat deposition
x
To
Beam of length L, thickness t, width w;
beam thermalized on one side at To
uniform heat deposition from one side, q (W/m2)
considering k constant with T
Boundary conditions: a) for x= 0 T=To (heat sink); b) q=0 for x=L (isolated tip)
Integrating and imposing the 2 boundary conditions:
x
T
T0
TL
L
practical interest:
calculate thickness
of a thermal shield

## Slide 30: Residual Gas Conduction
Viscous regime:
  At High gas pressure
  Classical conduction (q = - A k(T) dT/dx) with k independent of pressure
  but natural convection must be included
Molecular regime:
  At low gas pressure
  Kennard’s law
  Conduction is proportional to P
  Ω depends on gas species (for helium Ω = 2.13 W/m2.Pa.K)
  α(T)  accommodation coefficient depending on gas species, T1, T2 and surface geometry (applicable for flat parallel surfaces, coaxial cylinders and spheres)
d
T1
T2 (>T1)
λmolecule = mean free path
λmolecule << d  Viscous regime
λmolecule >> d  Molecular regime
Accommodation coefficient αi
Temp. [K] | Helium
300 | 0.3
80 | 0.4
20 | 0.6
4 | 1

## Slide 31
Thermal Radiation

## Slide 32: Thermal Radiation
Black surface: α = 1
Practical interest for cryostats shielding:
every gap acts as a black surface (example: 1 cm2 gap
exposed to a 293 K surface (e.g. vac.vessel with ɛ = 0.2)
receives ~10 mW

## Slide 33: Black body radiation
Total emissive power (integrating over λ):
In cryogenics: far infrared region:
(emissivity not necessarily related to surface appearance)
Emissive power (monochromatic)
In practice: a blackbody at 293 K emits ~ 420 W/m2:

## Slide 34: Radiation heat exchange between black bodies
Radiation from A1 to A2:
Radiation balance between A1 and A2:
Radiation from A2 to A1:
Note:

## Slide 35: Non-black surfaces: the diffuse-gray model (real surfaces)
Diffuse-gray emitter (good approximation for real surfaces)
Total hemispheric emissivity:
Similar considerations can be made for adsorptivity and reflectivity
The Diffuse-gray model:
  Gray
  A diffuse emitter, absorber and reflector
  Opaque (no transmittivity)
(Note: @ cryo temp. ε is strongly T dependent)

## Slide 36: Emissivity of various materials as a function of T
(R.B.Scott, Cryogenic Engineering, (Van Nostrand, New York, 1959; Y.S.Touloukian, Thermophysical Properties of Matter, (Plenum Press, New York, 1995))
0.1
Strong T dependence (quasi proportional to T)
Emissivity reduces with T
At cryogenic temperatures low emissivity in the far infrared is not necessarily related to surface brilliance

## Slide 37: Radiation between 2 diffuse-gray enclosures
Radiation balance between A1 and A2:
For 2 enclosed cylinders or spheres (not necessarily concentric!):
A1
A2
  Hints to reduce heat load in a cryostat:
  Reduce A2 (vac.vessel as small as possible)
  Small emissivities: ε1 reduced by low T; ε2 at RT & moderated by A1/A2

## Slide 38: Radiation between 2 diffuse-gray flat plates
Radiation balance between A1 and A2 (A1=A2=A):
T1
T2 (<T1)

## Slide 39: Radiation with an intermediate floating shield
Radiation balance between A1 and A2:
A1 to S gap
S to A2 gap
For flat surfaces approximation, and same ɛ, it becomes :
 ½ of the rate without shield (see previous slide)
  Hint: to reduce heat loads in a cryostat:
  - Add one (or more) intermediate shields

## Slide 40: Multi Layer Insulation (MLI)

## Slide 41: MLI principle
Low emissivity of aluminium layer
Multi-layer to enhance radiation protection:
  multi reflection of radiation…
Minimal thermal conductivity between reflective layers: interposing of isolating layers
  Reduced inter-layer thermal conduction heat loads
Enhanced performance @ low T  use actively cooled shield
  Lower emissivity of reflective material layers @ low T
  Reduce radiation from inner-most layers, cooled at T of shield
  Extract heat @ thermal shield T  more efficient heat extraction
Qc
Qr
Reflective film
Insulating spacer
Colder wall (actively cooled
thermal shield)
Warm wall
Tw
Tc < Tw

## Slide 42: MLI: How many reflective layers (N)?
Radiation vs. conduction, two conflicting phenomena
Radiation reduces as 1/N
Conduction is proportional to packing density (N/mm)
Packing density should be limited  typically ~ 25 N/cm
  Avoid “compressed” blankets, do not put as much MLI as possible…
  Do not forget space allocation for MLI blankets
  Consider differential thermal contractions wrt support (Al shields, cold mass…): blankets must remain loose at cold
Radiation
Total heat flow
Conduction
Arbitrary units
No. layers/cm

## Slide 43: Multi Layer Insulation (MLI): a simplified calculation model
A simplified model:
  Radiation reduction
  Solid conduction
N = No. of reflective layers
,  = average thermal conductivity
and emissivity constants of the MLI system
(obtained experimentally. For LHC cryostats:
=1.401 10-4, =3.741 10-9)
1
2
Qc
Qr
Reflective film
Insulating spacer
Considering the complexity of the phenomena
Involved, an experimental characterisation of MLI
performance, in particular for large machines, must
be made. However, abundant literature data available.

## Slide 44: LHC Multi Layer Insulation (MLI)
Features:
1 blanket (10 reflective layers) on cold masses (1.9 K)
2 blankets (15 reflective layers each) on Thermal Shields (50-65 K)
Reflective layer: double aluminized polyester film
Spacer: polyester net
Stitched Velcro™ fasteners for rapid mounting and quality closing
1 blanket on CM, 2 on thermal shield
Velcro™ fasteners
Blanket manufacturing
Velcro™ fasteners (inner side)

## Slide 45
Measured thermal performance on LHC

## Slide 46: LHeII calorimetric measurements of 1.9 K static heat loads in LHC
Transformation in p-T helium phase diagram during warm up
Schematic of a standard arc cell, a common superfluid helium bath of 106 m cooled by a unique heat exchanger tube.
C. Maglioni, V. Parma: “Assessment of static heat loads in the LHC arc, from the commissioning of sector 7-8”, LHC Project Note 409, 2008.
Average heat load to cold mass (10 MLI layers) ~ 0.2 W/m
Rescaled on cold mass surface and subtracting solid conduction contributions (lab tests on components) :
10 layers MLI between 50K and 1.9 K  ~ 0.054 W/m2
Static HL natural warm-up of cryogenic subsector after stop in cooling
 Practical figure: 50 mW/m2
Change of internal energy
Length of string of magnets

## Slide 47: Non-isothermal cooling of LHC thermal shield (2’700 m)
Average heat load @ 50K of ~ 4 W/m
Thermal shielding with MLI (30 layers)
Rescaled on cold mass surface and subtracting solid conduction contributions (lab tests on components):
30 layers MLI between 300K and 50 K  ~ 1 W/m2
Helium flow (measured)
Δh
Thermal shield static heat load profile along sector 7-8
LTT
Specific enthalpy change
(T measurements and He properties)
[W/m]
Distance between T sensors
 Practical figure: 1 W/m2

## Slide 48
Numerical application on the LHC Cryostat

## Slide 49: Application to an LHC-like cryostat
heat loads HL will be calculated for a 1-m cryostat unit length
Vacuum vessel diameter: 1m (AVV= π x 1= 3.14 m2)
Cold mass diameter: 0.6 m (ACM= π x 0.6 = 1.88 m2)
T cold mass: 2 K
T vac.vessel: 293 K
Budgets: HLCM ~ 0.2 W/m; HLTS ~ 5 W/m

## Slide 50: a) Bare cold mass
Emissivity cold mass: εCM = 0.12
Emissivity vac.vessel: εVV = 0.2
HLCM = 63 W
Budget for LHC is ~0.2 W  HL too high

## Slide 51: b) Cold mass wrapped with 1 layer of Al foil
Emissivity of Al foil (at 2 K): εCM = 0.06
HLCM = 40 W
 HL still too high

## Slide 52: c) Cold mass wrapped with 30 layers of MLI
HL from 290 K with 30 MLI layers (calculated with MLI formula): 1.2 W/m2
HLCM = 1.2x1.88= 2.3 W
 HL still 1 order of magnitude too high

## Slide 53: d) Addition of thermal shield actively cooled
Thermal shield diameter: 0.8 m (ATH= = π x 0.8 = 2.51 m2)
Thermal shield at intermediate T  80 K
Emissivity of Al (at 80 K): εTS = 0.1
HLCM = 0.26 W
HLTS = 79 W
 Close to budget
 too high (Budget for LHC is 5 W)

## Slide 54: e) Wrapping of MLI around thermal shield
HL from 290 K with 30 MLI layers 1.2 W/m2
HLTS = 1.2x2.51= 3.01 W
Within budget for LHC (5 W)
HLCM = 0.26 W
 Close to budget

## Slide 55: f) Adding 1 Al foil around cold mass
Emissivity of Al foil (at 2 K): εCM = 0.06
HLTS = 3.01 W
Within budget (5 W)
HLCM = 0.18 W
 within budget (0.2 W)

## Slide 56: g) What in case of bad vacuum (He leaks)?
 Residual gas molecular conduction:
From table, for He at 2 K: αCM=1, for He at 80K αTH=0.4  α = 0.47
for He, Ω = 2.13 W/m2.Pa.K
For 1 Al foil on cold mass, in case of degraded vacuum:
For P = 10-3 Pa (10-4 mbar):
(still quite good vacuum)
For P = 10-1 Pa (10-2 mbar):
(degraded vacuum)
Exceeds budget
2 orders or magnitude higher than
budget!!

## Slide 57: h) Add MLI around the cold mass
10 MLI layers on cold mass
Using measured data
  In good vacuum (<10-3Pa): 50 mW/m2
  Under degraded vacuum (~10-1 Pa): ~2W/m2
HLCM = 1.88x0.05 = 0.09 W
(HLTS = 3.59 W)
HLCM = 1.88x2 = 3.8 W
MLI cuts residual conduction by 4 !!
HL even lower
Important note: MLI on helium vessels also necessary to reduce by about 7 condensation heat fluxes in case of accidental cryostat venting with air (bare surface: q ~4 W/cm2 ; 10 layers of MLI: q ~0.6 W/cm2 )

## Slide 58: Summarizing
Case | 2K heat loads | 80 K heat loads
a) Bare Cold mass | 63 W | N.A.
b) Cold mass with 1 Al foil | 40 W | N.A.
c) Cold mass with 30 MLI layers | 2.3 W | N.A.
d) 1 thermal shield at 80K, no MLI | 0.26 W | 79 W
e) 30 MLI layers on thermal shield | 0.26 W | 3.01 W
f) As e) + 1 Al foil on cold mass | 0.18 W | 3.01 W
g) As f) but degraded vacuum | up to 15 W (10-1 Pa) | > 3.01 W
h) +10 MLI layers on cold mass | 0.09 W in good vac. 3.5 W in deg.vac. | 3.59 W > 3.59

## Slide 59: What about a second actively cooled shield?
70 K cooling
5-20 K cooling
Experimental program for LHC cryostat in the late nineties (Cryostat Thermal Model, CTM)
A 20 K active cooled screen with 10 MLI layers
An estimated saving of up to ~0.15 W/m at 1.9 K
but the an equivalent increase at the 5-20 K level (~5 times less costly)
Overall electrical power saving: ~ 100 Wel/m
Additional hardware (line, MLI, supports,etc)  higher capital cost
Additional assembly complexity
Breakeven only after ~10 years of operation
For LHC we decided to keep 1 active shield at 70K
The Cryostat Thermal Model

## Slide 60: Thermal shield: what thickness ?
Aluminium shield, in Al 5052
Actively cooled by 1 cryo line at 80K
 Average conductivity: k = 80 W K-1m-1
Uniform heat deposition:
HLTS= 3.59 W  q = 3.59/(0.8x1x π) = 1.43 W/m2
Calculate thickness with the requirement:
Azimuthally quasi iso-thermal shield:
  ΔTmax= Tmax-Tmin ≈ 5 K
Remembering the formula yielding ΔTmax:
Replacing L by ½ circumference of diameter D (Tmax opposite to cryo line):
Tmin
Tmax
Cooling
line
(for LHC, 2.5mm thick Al 1100 equivalent)

## Slide 61: LHC thermal shields
Aluminium alloy 6063 extrusions and 1100 top sheets
Al 6063 extrusions
LHC Main Cryostat
(Cross-Section)
Al 1100 sheets
He cooling (50-65K)

## Slide 62
Heat transfer to cryogenic fluids

## Slide 63: 2 main mechanisms of interest for cryostats
Vaporisation in pool boiling (2-phase)
  Latent Heat (LH) of vaporisation
  Isothermal cooling (T constant if P constant)
Vapor
U
Tin
Tout
Cryogen | Latent Heat (at 1tm) [kJ/kg] | [mg/s] | [l/h] (liquid) | [l/min] (gas NTP)
Helium | 21 | 48 | 1.38 | 16.4
Nitrogen | 199 | 5 | 0.02 | 0.24
Vaporisation under 1 W heat load
Forced internal (tube) convection of single-phase fluid:
  Non-isothermal cooling: enthalpy change of fluid
  Depends from thermo-hydraulics of the flow (see next slide)
  Used in cooling of thermal shields (supercritical He)
For more cooling mechanisms see dedicated lectures

## Slide 64: Forced Convection Heat Transfer
D
L
U
Forced flow of coolant fluid in round tube cooling lines
Considering hydro-dynamically and thermally fully developed flow
Uniform wall heat flux (linear T profiles)
Convection heat transfer from wall to fluid:
Reynolds No.:
Nusselt No.:
For laminar flow:
For turbulent flow, NuD=f(ReD, Pr):
Case of a Thermal
Shield
ReD>2000  turbulent flow,
ReD<2000  laminar flow
Enthalpy balance along the line L:

## Slide 65: Frictional pressure drop in a tube
Pressure drop along tube

## Slide 66: Cryogenics considerations

## Slide 67: Helium as a coolant
Property | Units | 4He | N2
Boiling T (at 1 atm) | K | 4.2 | 77.3
Critical temperature | K | 5.2 | 126.1
Critical pressure | 105 Pa | 2.23 | 33.1
Latent Heat of evaporation (at 1 atm) | kJ/kg | 21 | 199
Enthalpy between T boiling and 300K | kJ/kg | 1550 | 233
Liquid density (boiling at 1 atm) | kg/m3 | 125 | 810
Saturated vapor density (at 1 atm) | kg/m3 | 17 | 4.5
Gas density (at 1 atm 273.15K) | kg/m3 | 0.18 | 1.25
Liquid viscosity (at boiling T) | μPa.s | 20 | 17
Pressurized He II, Magnets: LHC, Tore Supra
Pressurized He I, Magnets, HERA, Tevatron
Saturated He II, SRF: CEBAF, TTF, SNS, EXFEL, ESS, ILC
Pool boiling He I, SRF: HERA, LEP, KEKB
Supercritical helium: cooling of thermal shielding

## Slide 68: Refrigeration efficiency (Carnot principle)
A refrigerator extracts a heat flow at a temperature below ambient and rejects it at a higher temperature (normally ambient)
The Carnot cycle defines the minimum mechanical work/power (i.e. Maximum Coefficient of Performance, COP) which depends only on Tw and Tc
All real machines have a lower efficiency (irreversibilities), expressed in fraction of COP
Fluid | T [K] | Carnot factor (W/Qc) [W/W] (considering Tw=293K)
LN2 | 77 | 2.8
LH2 | 20.4 | 13.4
LHe | 4.2 | 68.4
LHe | 1.8 | 161.8
Tc
Tw
Work (W)
Qc
Qw
T
S
reversible adiabatic (isoentropic)
reversible isothermal
Carnot cycle
Tw
Tc
W

## Slide 69: Efficiency for large cryoplants
State-of-the-art figures for large cryo-plants (LHC-like, ~18 kW @ 4.5K):
COP @ 2 K  ~ 15% Carnot (990 Wel/Wth)
COP @ 4.5 K  ~ 30% Carnot (210 Wel/Wth)
COP @ 50 K  ~ 30% Carnot (16 Wel/Wth)

## Slide 70: He Liquefaction: conversion to 4.5K isothermal capacity
He refrigeration He Liquefaction
Wliq =T0 Δscond–Qcond + T0 ΔSprecool – Qprecool = T0 ΔScond– ΔHcond+ T0 ΔSprecool – ΔHprecool
For T0 = 300 K, and S, H tables  Wliq = 6,600 W per 1 g/s of He liquef.
Take 66 Wel/Wth as minimum specific refrigeration work @ 4.5 K (Carnot):
 1 g/s of liquefied He (6,600 W) is equivalent to ~100 Wth @ 4.5 K
Refrigeration: recondensing cold vapours
  Limited use of cryogenic power at boil-off (latent heat of evaporation)
  Simpler cryoplant
Liquefaction: precooling + recondensing cold vapours
  Availability of cold vapours enthalpy up to RT
  Added complexity
(combining 1st & 2nd principle of thermodynamics, and introducing entropy)
W condensation
W Pre-cooling
W =T0 [Qc/Tc – Qc] = T0 ΔSc – Qc

## Slide 71: Thermally efficiency solid conduction: heat intercepts, helium vapour cooling

## Slide 72: Solid conduction in cryostats
LHC Main dipole
Magnet Support Posts (GFRE)
Solid conduction paths:
Supporting systems
Current leads
RF main coupler
Beam tubes Cold-to-Warm (CWT) transitions
Instrumentation feed-throughs
Necks (vertical cryostats)
Current leads
Neck
LHC instrumentation capillary at assembly
CWT SPL cryomodule

## Slide 73: Solid conduction and heat intercepts
L
A
Tw=293K
Tc=2K
simple solid conduction
L
A
Q @ 80K
L1
L
A
Q @ 80K
Q @ 8K
L1
L2
Minimizing using cost factors:
C1 = 16 w/w
C2 = 210 w/w
C3 = 990 w/w
1 heat intercept at optimal distance
2 heat intercepts at optimal distance
 L1
 L1, L2

## Slide 74: LHC supports
Two intercepts, 5K & 75K
One intercept, 75K
No intercepts
5-10 K heat intercept
GFRE composite
column
Vacuum vessel
Interface flange: 293 K
Cold mass interface flange: 2 K
50-65 K heat intercept
4-mm thickness, single-part composite column (integrating interface flanges)
Manufactured by Resin Transfer Moulding (RTM):
  Suited to a large-scale industrial production (4’700 units)
  High reproducibility in thermo-mechanical properties

## Slide 75: Vapour cooling in solid conduction
x, T(x)
@ Tl
Tw
L
A
Large enthalpy in He vapours (1550 kJ/kg from 4.2K to 300K)  usable cooling capacity
If , which is the residual heat to the bath, is equivalent to the evaporation (i.e. self-sustained):
Vapour cooled wall
Assuming perfect exchange (T gas = T wall)
attenuation factor (w.r.t. solid conduction)
Reduced heat conduction in self-sustained helium cooling for selected technical materials

## Slide 76: Vapour cooled RF coupler for SPL
RF couplers with He gas cooled double walled tube
When RF is on, a distributed vapour cooling is essential to contain distributed RF heating (local heat intercepting can hardly provide efficient cooling)

## Slide 77: Vapour cooling of current leads
x, T(x)
@ Tl
Tw
L
A
With:
I, current in the lead
Cp(T), specific heat
f = 0  no cooling
f = 1  perfect heat exchange (T(x) lead=T(x) vapour)
x+dx, T(x+dx)
conduction
cooling
(efficiency 0 < f < 1)
resistance heating
For current lead material following the Wiedmann-Franz law (most metals and allows, Cu for example):
ρ(T) and k(T) are correlated! (good electrical conductors are also good thermal conductors)
Minimising heat in-leaks is independent of material choice for normal conducting materials
Substituting in the above equation and integrating it for variable f efficiencies…(next slide)

## Slide 78: Heat load to bath per unit current
No cooling
47 W/kA
Perfect cooling
1.04 W/kA
Enhancing thermal performance can be achieved with materials which do not follow the WF law
High Temperature Superconductors, for example, have zero resistivity and are relatively bad thermal conductors up to high temperatures.  more in specific lecture
Example of off-the-shelf lead from AMI
(current rating up to 10 kA)

## Slide 79: End of 1st Part…

## Slide 80: Insulation vacuum and construction aspects

## Slide 81: Leaks
Units:
A leak is a throughput, normally given symbol qL
Common units are:
  mbar.l/s atm.cc/s torr.l/s Pa.m3/s (SI unit)
  With a leak rate of 1 mbar.l/s a volume of 1 litre will change in pressure by 1 mbar in 1 second
  Units of mbar.l/s equivalent to atm.cc/s
    Eg immersed in water:
    A leak of 1 atm.cc/s would produce a bubble of 1 cm3/s
    A leak of 10-3 atm.cc/s would produce a bubble of 1 mm3/s
Flux through a leak will be different depending on the prevailing conditions (temperature, pressure, gas type)
Unless otherwise stated, a ‘standard helium leak rate’ in mbar.l/s implies:
  Helium as tracer gas,
  Under vacuum test,
  Helium at 1 barabs and 100% concentration
  System at 20 °C.

## Slide 82: Leak tightness
No vacuum vessel is leak-tight, nor need it be
Define the satisfactory leak rate needed to remain within the needed pressure in a given time:
Normally 2 sources of pressure increase: leaks and outgassing
leak +
outgassing
leak
outgassing

## Slide 83: Cryopumping: cryo-condensation
2-4.5 K surfaces have high pumping speed & capacity
Gas species have very low saturation vapour pressure except helium
Without helium leaks, equilibrium pressures < 10-6 mbar are obtained
 helium leaks are the real issue

## Slide 84: MLI outgassing
Outgassing in cryostats is normally dominated by MLI outgassing
For an LHC insulation vacuum sector (80 m3, 250 m2/m length of MLI, 214m length) exposed to ambient air for several weeks, we obtain ~ 1 e-3 mbar at RT after ~ 200 hrs pumping (S = 100 l/s). Equivalent to ~ 2 e-10 mbar.l/s/cm2 of MLI.

## Slide 85: Insulation vacuum and heat loads from residual gas conduction
LHC dipole Cryostat Test Model results
Insulation vacuum is about minimising heat transfer due to residual gas conduction (wrt to radiative and conductive heat transfer)
Only helium leaks can degrade the vacuum
Determine the maximum acceptable degraded helium pressure for cryostat (only helium matters for LHC)
  For LHC  ~ 10-4 mbar (10-2 Pa)

## Slide 86: LHC Strategy
Consider no mechanical pumping on insulation vacuum during operation
Determine maximum helium degraded vacuum pressure  10-2 Pa
Determine helium cryosorption capacity of cold surfaces
  For LHC ~ 100 mbar.l of helium @ 1 e-4 mbar per 214 m of cryostat
Determine the leak rate that will saturate the cold surfaces of 1vacuum sector after 200 days of LHC operation  < 5 10-6 mbar.l/s
Use fixed turbos during thermal cycles and as ‘backup’ in case the tightness specification cannot be immediately reached
Apply cold/warm correlation for leak rates (considered to be up to x1000 at cold)  < 5 10-9 mbar.l/s per vacuum sector
Allocate higher levels of leak tightness to sub-assemblies and components in one same vacuum sector  down to < 10-11 mbar.l/s
 These very low levels of leak tightness on steel metal work and welded piping assemblies are extremely challenging for construction and testing, especially for large industrial productions (e.g. LHC)

## Slide 87: Leak detection - Common methods
TEST METHOD
Bubble test
Pressure
variation
Sniffing
halogens or H2N2
Helium mass
spectrometer
102
101
100
10-1
10-2
10-4
10-5
10-6
10-8
10-10
Flux in atm.cm3/s or mbar.l/s
10-12
10-3
Under vacuum
Over pressure
Over pressure
Over pressure (sniffing)
Under vacuum
Over pressure
Residual gas
analyser
Under vacuum
Courtesy of P.Cruikshank

## Slide 88: He leak detection methods
Vacuum
He
Leak Detector
Helium
pistol
Part
to test
helium
Q
q = helium flux in mbar. l/s
He
5 bar
He
Leak Detector
Sniffer
in helium
cloud
Courtesy of P.Cruikshank
Sniffing
(leak localisation)
Vacuum
(leak detection)

## Slide 89: Leak detection with clam-shell
Clam-shells for pipe geometries for LHC
A practical detection method for
circumferential welds
Pumping of reduced volumes
Particularly interesting for helium polluted circuits (e.g. magnets cold tested in helium)

## Slide 90: Welding
Tungsten Inert Gas (TIG) and Metal Inert Gas Welding (MIG) are the most commonly used processes in cryostat fabrication
Qualification of welding procedures and personnel required to fulfil mechanical requirements imposed by pressure vessel codes
Design of welded seams must be carefully chosen to avoid sources of impurities and defects
Non-destructive testing is often necessary
Tungsten inert gas welding
Metal inert gas welding

## Slide 91: Design of vacuum facing welds

## Slide 92: Design of pressure bearing welds
EN 13445-3 annex A is a good reference for designing pressure bearing welds. EN 1708-1 is also a very useful hamonised standard. Some examples:
Longitudinal welds
Flat ends
Nozzles
Circular welds

## Slide 93: Welders and procedures qualification
 | Steel | Aluminium
Welding procedure approval | EN ISO 15614-1:2004 Specification and qualification of welding procedures for metallic materials - Welding procedure test - Arc and gas welding of steels and arc welding of nickel and nickel alloys | EN ISO 15614-2:2005 Specification and qualification of welding procedures for metallic materials - Welding procedure test - Arc welding of aluminium and its alloys
Qualification of welders | EN 287-1:2004 Qualification test of welders - Fusion welding - Steels | EN ISO 9606-2:2004 Qualification test of welders - Fusion welding - Aluminium and aluminium alloys
Qualification of welding operators | EN 1418:1998 Welding personnel - Approval testing of welding operators for fusion welding and resistance weld setters for fully mechanized and automatic welding of metallic materials | 

## Slide 94: Some typical mistakes causing leaks
Flexible hose
Same weld including braid/ring/bellows  separate weld functions
Flanges made of cold rolled material
Leak through material inclusions  QA of raw materials, or 3-D forged flanges (but expensive!)
Sheet metal work vessel
Leak through material crack  QA of raw materials,

## Slide 95: Brazing
Often the only solution to join different materials (ex: copper to stainless steel; stainless steel to ceramics…)
Vacuum brazing (no flux required) gives the most reliable joints, but at a cost
Thorough cleaning after brazing with flux is mandatory. Poor cleaning often results in the development of leaks in stainless steel due to corrosion!
Useful standards for brazing specification and execution:
Example of flame brazed stainless steel to copper transition for a thermal shield cooling circuit
EN 13134:2000 Brazing - Procedure approval
EN 13133:2000 Brazing - Brazer approval
EN 12797:2000 Brazing - Destructive tests of brazed joints
EN 12799:2000 Brazing - Non-destructive examination of brazed joints
EN ISO 18279:2003 Brazing - Imperfections in brazed joints

## Slide 96
Mechanical considerations

## Slide 97: Thin shells under internal pressure

## Slide 98: Radial buckling under external pressure
Non-linear phenomenon. Actual critical pressure depends on initial imperfections: Safety factor needed!
Buckling critical pressure for a thin tube of infinite length
A conservative rule of thumb for stainless steel tubes under vacuum:
If we use a safety factor of 3:
Alternatively, we need to add reinforcements
Example:
r = 500 mm
t > 18.5 mm

## Slide 99: LHC dipole Vacuum Vessels
Main features:
Pipeline standard size: 36-inch OD (1013 mm), 12-mm thick, low carbon steel (DIN GS-21 Mn5) tubes
St. steel extremity flanges
Material resilience: > 28 J/cm2 at -70ºC
Forged cradles, welded rings reinforcements
Dimensional stability:
  Stress relieving
  Final machining to achieve tolerances at interface
3-D dimensional control in Industry
Out-doors storage at CERN
Final lathe machining
Forged cradle
Stress relieving
Production:
1250 units
2 firms
4 yrs of production

## Slide 100: Pressure vessel codes regulations
Category | Conf. assessment module | Comment
SEP | None | The equipment must be designed and manufactured in accordance with sound engineering practice. No CE marking and no involvement of notified body.
I | A | CE marking with no notified body involvement, self-certifying.
II | A1 | The notified body will perform unexpected visits and monitor final assessment.
III | B1+F | The notified body is required to approve the design, examine and test the vessel.
IV | G | Even further involvement of the notified body.
For vessels with non-dangeroureous gases (cryogenic liquids are treated as gas)
Pressure European Directive 97/23/EC (PED) is obligatory throughout the EU since 2002
  Applies to internal pressure ≥ 0.5 bar
  Vessels must be designed, fabricated and tested according to the essential requirements of Annex 1 (Design, safety accessories, materials, manufacturing, testing, etc.)
  Establishes the conformity assessment procedure depending on the vessel category, which depends on the stored energy, expressed as Pressure x Volume in bar.L

## Slide 101: Harmonised codes and standards
Harmonised standards give presumption of comformity with the PED, within their scope. Uselful codes for cryostat design and fabrication:
  EN 13458-1:2002 Cryogenic vessels - Static vacuum insulated vessels - Part 1: Fundamental requirements
  EN 13458-2:2002 Cryogenic vessels - Static vacuum insulated vessels - Part 2: Design, fabrication, inspection and testing + EN 13458-2:2002/AC:2006
  EN 13458-3:2003 Cryogenic vessels - Static vacuum insulated vessels - Part 3: Operational requirements + EN 13458-3:2003/A1:2005
  EN 13445-1:2009 Unfired pressure vessels - Part 1: General
  EN 13445-2:2009 Unfired pressure vessels - Part 2: Materials
  EN 13445-3:2009 Unfired pressure vessels - Part 3: Design
  EN 13445-4:2009 Unfired pressure vessels - Part 4: Fabrication
  EN 13445-5:2009 Unfired pressure vessels - Part 5: Inspection and testing
  EN 13445-8:2009 Unfired pressure vessels - Part 8: Additional requirements for pressure vessels of aluminium and aluminium alloys
Other codes such as the French CODAP or the American ASME Boiler and Pressure Vessel Code can be used, but proof of conformity is at the charge of the manufacturer

## Slide 102: Useful material standards for cryostats
Plates and sheets | • EN 10028-1:2007+A1:2009 Flat products made of steels for pressure purposes - Part 1: General requirements • EN 10028-3:2009 Flat products made of steels for pressure purposes - Part 3: Weldable fine grain steels, normalized • EN 10028-7:2007 Flat products made of steels for pressure purposes - Part 7: Stainless steels
Tubes | • EN 10216-5:2004 Seamless steel tubes for pressure purposes - Technical delivery conditions - Part 5: Stainless steel tubes • EN 10217-7:2005 Welded steel tubes for pressure purposes - Technical delivery conditions - Part 7: Stainless steel tubes
Forged blanks | • EN 10222-1:1998 Steel forgings for pressure purposes - Part 1: General requirements for open die forgings • EN 10222-5:1999 Steel forgings for pressure purposes - Part 5: Martensitic, austenitic and austenitic-ferritic stainless steels
Castings | • EN 10213:2007 Steel castings for pressure purposes
Pipe fittings | • EN 10253-4:2008 Butt-welding pipe fittings - Part 4: Wrought austenitic and austenitic-ferritic (duplex) stainless steels with specific inspection requirement
Bars | • EN 10272:2007 Stainless steel bars for pressure purposes
Aluminium | • EN 12392:2000 Aluminium and aluminium alloys – Wrought products – Special requirements for products intended for the production of pressure equipment (choose materials included in the list given in EN 13445-8 section 5.6)

## Slide 103: Design stresses for some materials
Material | Rp1.0 (MPa) | f (MPa) | ftest (MPa)
1.4306 (304L) | 240 | 160 | 228
1.4435/1.4404 (316L) | 260 | 173 | 247
1.4406/1.4429 (316LN) | 320 | 213 | 304
AW 5083-O/H111 |  | 83 | 
Design stresses for plates less than 12 mm thick applicable to membrane stress (safety factor 1.5 included) according to EN 13445-3
For stainless steels:
For aluminium-magnesium alloys:
Material | Rp1.0 /Rm (MPa) | f (MPa) | ftest (MPa)
AW 5083-O/H111 | 125/270 | 83 | 119

## Slide 104: Best practices
Using a coherent set of standards throughout the lifecycle of the cryostat is the simplest and safest approach. As an example when using only EN harmonised standards:
  Error margins of pressure relief devices are taken into account in the design rules
  The design rules are only applicable if the material has enough ductility
  Materials certified for pressure vessels have measured minimum fracture toughness
  Safety factors included in buckling formulae take into account shape imperfections up to the allowable tolerances layed out in the manufacturing section of the standards
  The extent of welding inspection must be compatible with the joint coefficient used in thickness calculations
  Coherence of test pressure and testing procedure with the design rules

## Slide 105
Thermo-mechanical considerations

## Slide 106: Thermal expansion of some metals

## Slide 107: Thermal expansion of some plastics

## Slide 108: Thermal stress: 3 cases
ΔT
A) Restrained component
B) Assembly of different materials
Material 1: α1, E1, A1
Material 2: α2, E2, A2
restrain
ΔT
C) Different cooling ΔT1(t) ≠ ΔT2(t) (different material diffusivity or different cooling)
Material 1: α1, E1, A1
Material 2: α2, E2, A2
restrain
ΔT2(t)
ΔT1(t)

## Slide 109
Supporting systems

## Slide 110: Supporting system
Mechanical housing of cryogenic devices (supporting systems):
  Supporting of (sometimes heavy) devices
  Accurate & reproducible positioning (almost always)
  Precise alignment capabilities (SC devices in accelerators)
Many solutions available:
  Tie rods
  Suspended posts
  Compression posts
  …other
Each having specific advantages/drawbacks depending on:
  SC device’s weight and cryostat assembly methods
  Vacuum vessel external supporting (supported? Suspended?)
  Adjustment of cold mass inside vacuum vessel
  …
For the LHC, the compression posts were preferred because of :
  Heavy cold masses (~30 tons!)  supported on jacks on tunnel floor
  Cryostat assembly based on sliding (or rolling through) of cold mass standing on supports
  No need for adjustment, magnets individually fiducialised and machine aligned w.r.t. external cryostat-mounted fiducials

## Slide 111: LHC dipole cryostat assembly
Pulling through sliding on vacuum vessel
Assembly bench

## Slide 112: LHC Supporting system
The design is a trade-off between 2 conflicting requirements:
High flexural stiffness (for mechanical stability)  thick and bulky structure
Low heat in-leaks  thin and slender structure and low conductivity material
 Flexural stiffness/conductivity is an interesting figure of merit in the choice of the material
x
q
F
Q
(The acceptable x and θ
are alignment requirements)
(F is mainly the result of
interconnect forces
and gravity component)
(Q is to be within budget)

## Slide 113: Choice of the material
St.steel  interesting below 20K
G10 and Ultem 2300  preferable at 20K < T < 300K
Other interesting material: Carbon-fiber Epoxy  also interesting below 20 K (not shown in diagram)
For LHC, a Glass-fiber Epoxy Composite (GFRE) was chosen:
Good conductivity/flexural stiffness
Widely available on the market  cost effective for large production (5000 units!)
…but a specific thermal conductivity validation campaign was needed.

## Slide 114: LHC supporting system
No. of supports, spacing and positions:
  2 support posts whenever possible:
    Isostatic: well known forces on cold mass/supports/vacuum vessel, not conditioned by handling
    Optimise spacing to minimize vertical sag
  Add 3rd support post if necessary for long cold masses:
    Limit vertical sag to acceptable values (cold mass straightness)
    Hyper-static: precautions when handling, use of specific girders
  Position of support posts on vacuum vessel:
    Always above the external jacks  direct load transfer from cold mass to ground, hence the vacuum vessel is unstressed (only vacuum loads).

## Slide 115: The LHC vacuum vessel, a 3 supports solution

## Slide 116: Longitudinal thermal contractions
Cold mass, thermal shield, support posts and vacuum vessel must be free with each other to cope with longitudinal thermal contractions
One fixed point per each component
Leave plays to cope with all extreme T cases (ex. Cold mass cold, thermal shield warm)
Guided sliding of cold mass onto vacuum vessel
Flexible thermalisations anchors
(example of an LHC special quadupole)
Fixed post
Sliding posts
Δl cold mass= 30 mm (10m x 3mm/m)
Δl thermal shield= 40 mm (10m x 4 mm/m)

## Slide 117: Other solutions
HERA Dipole

## Slide 118: Other solutions
RHIC Dipole

## Slide 119
Other solutions
Fixed support
Sliding support
Inertia beam
Invar longitudinal positioner
External supports (jacks)
RF coupler (with bellows)
Tesla/TTF/ILC cryomodule

## Slide 120: Other solutions
SNS high beta cryomodule
“Space-frame” and tie rods

## Slide 121: Other solutions
Inter-cavity support ensures alignment
Double tube of the power coupler as main support
Vacuum vessel
SPL
cryomodule

## Slide 122: Pressure relief protection systems

## Slide 123: Pressure relief protection systems
Cryostats include large cold surfaces, inventory of cryogenic fluids, sometimes large stored energy (e.g. energized magnets)
  a potentially unstable energy storage which will tend to find a more stable state of equilibrium
  Through a thermodynamic transformation which can be sudden and uncontrolled with a dangerous increase of pressure
Protect personnel (burns, ODH) and equipment (direct and collateral damage)
Risk hazards:
  Sources of pressure:
    Compressors connected to cryo lines
    Connection to higher pressure source (e.g. HP bottles)
    Heating of “trapped” volumes (typically in a circuit between valves) during warm-ups
    Helium leak to insulation vacuum, with consequent increased conduction/convection heat loads to cryogenic liquid vessels
    Cryo-condensed air leaks on cold surfaces and consequent pressure increase and increased conduct/convection heat loads during warm-ups
    Heating/vaporization of cryogens from sudden release of stored energy in SC device (e.g. quench or arcing in a SC magnet circuit)
    Uncontrolled air/nitrogen venting of insulation vacuum with sudden condensation on cold surfaces
    Uncontrolled release of cryogenic fluid to higher T surfaces (thermal shield and vacuum vessel), and consequent pressure increase and increased of conduction/convection heat loads to cold surfaces
Often the most critical

## Slide 124
A typical example…LHC 19th sept.2008
…5 MW arc in an interconnect
…~50 t of longitudinal;pressure force, ~50 magnets displaced …
…uprooting of jacks
~2 t of helium in first 2 minutes, estimated He peak mass flow of 20 kg/s
Insulation vacuum relief device

## Slide 125: General approach
Risk analysis & mitigation:
Make a thorough risk analysis and evaluate risk hazards
Identify mitigation measures (e.g. protections of exposed bellows and flanged connections)
Identify severity of consequences and appreciate probability of the event
Define the maximum credible incident(s) and design the safety relief system accordingly
The safety relief system must be designed to keep pressure rise within the limits of the Maximum Allowable Working Pressure (MAWP)
Design steps:
Estimate the heat exchange and its conversion to mass flow rates to be discharged
Check the sizing of piping (generally designed for normal operation) to the relief device and increase if necessary
Choose the type of safety device (burst disks, valves, plates) and size the safety device (DN and set pressure). Make use of safety device manufacturers formulas and charts
Size recovery piping downstream of safety device and check venting needs in the buildings where the release occurs (ODH issue)

## Slide 126: Pressure Safety Relief Devices
Cryogenic fluid vessel
  Typical ΔPmax < PS
  PS depends on the device (~few bar for SC cavities, up to ~ 20 bar for magnets)
  Define DN of valve and set pressure, PT
Vacuum
Air
Cryogenics fluid
Vacuum
Air
Cryogenics fluid
  According to European directive 97/23/EC and EN 13648 “Safety devices for protection
  against excessive pressure”
Vacuum vessel
  Typical ΔPmax < PS (0.5 bar relative to atm. for vac.vessels)
  Define DN of valve and set pressure, PT

## Slide 127: Cryogenic fluid vessel
The cryogenic fluid volume must be protected against over-pressure consecutive to unexpected heat transfers
Hazard: breach in insulation vacuum:
Uncontrolled air/nitrogen venting of insulation vacuum with sudden condensation on cold surfaces
Heat flux:
From 3 experimental sources internationally recognised:
  W. Lehman and G. Zahn, “Safety Aspects for LHe Cryostats and LHe Transport Containers,” ICEC7, London, 1978
  G. Cavallari, et. al., “Pressure Protection against Vacuum Failures on the Cryostats for LEP SC Cavities,” 4th Workshop on RF Superconductivity, Tsukuba, Japan, 14-18 August, 1989
  M. Wiseman, et. al., “Loss of Cavity Vacuum Experiment at CEBAF,” Advances in Cryogenic Engineering, Vol. 39, 1994, pg. 997.
Experimental values:
0.6 W/cm2 for a superinsulated tank of a bath cryostat
Up to 4 W/cm2 for a bare surface tank of a bath cryostat
Air
Cryogenic fluid
P↗

## Slide 128: Cryogenic fluid vessel (cont.d)
The safety device should be designed to relieve a mass flow equivalent to the highest heat load
Calculate the mass flow, Qm to be released by the safety device (EN13468-3.4)
2 cases for calculating mass flow Qm:
Below critical pressure (p<2.23 bar for helium):
  Bi-phase with liquid boil-off  take Lv (latent heat)
Above critical pressure (often the case):
  Supercritical fluid expelled  use a “pseudo latent heat” Lv’
P0 : relieving pressure [bara]
Pc : critical pressure [bara] (2.23 for He)
Qm : mass flow in [kg.s-1]
W : heat load [W]
L : latent heat in relieving conditions [J.kg-1] (20.103 at 1 bar for He)
vg/vl : specific volume of saturated gas/liquid at P0 [m3.kg-1]
L’ : specific heat input, see EN13468-3.4
h : enthalpy of the fluid [J/kg]
v : specific volume [m3.kg-1]

## Slide 129: Cryogenic fluid vessel (cont.d)
The minimum required flow area A is calculated with conservative assumptions on fluid properties
For compressible fluids, the mass flow through a restriction depends on the downstream pressure until a fixed Pb/P0 ratio (0.49 for helium)
= 0.49 for He

## Slide 130: Vacuum vessel
The vacuum vessel safety device is designed to relieve a mass flow equal to the highest incoming flow but at warmer temperature while keeping the vessel pressure within the PS
Identify the worst case scenario (highest mass flow and coldest fluid)
Often the worst case corresponds to a rupture of a cryogenic circuit:
    The cryogenic fluid flows into the vacuum vessel  the fluid vaporizes/expands in contact with the warm walls  the internal pressure increases until the safety device set pressure  the device opens and the fluid is relieved to atmosphere
Calculate the mass flow from the reservoir to the vacuum vessel
  Estimate the area of the breach in the cryogenic circuit
  Calculate the mass flow through an orifice
Pvv > Patm
Air at Patm
Cryogenics fluid

## Slide 131: Vacuum vessel (cont.d)
Calculate the minimum required flow area, A for the safety device
  Mass flow through the safety device = mass flow to the vacuum vessel
    Qm1=Qm2
    A > than the orifice area as Pb/P0 is lower and the gas is warmer.
  The flow area is highly dependent on the relief temperature, usually difficult to estimate
    First case Trelief = 300K
    If the device is too big, investigations are needed to estimate Trelief

## Slide 132: Examples of safety devices
  Some suppliers:
  Herose, Rembe, Ramseyer, Leser…
Safety valves
Burst disks
LHC pressure release plates
(DN200)

## Slide 133: Summary
Since Dewar’s invention, cryostats have evolved from simple containers for cryogens to sophisticated mechanical assemblies for SC accelerator devices for fundamental science as well as for industrial applications (e.g. NMR machines)
Though the understanding of the heat transfer phenomena involved in a cryostat have considerably progressed since the time of Dewar, the main outstanding innovation was the introduction of MLI, in the 50ties…
…But the enabling technologies, have greatly evolved from “simple” “glass-blowing” to covering a wide range of disciplines, enhancing performance of modern cryostats:
  Low thermal conductivity composite materials
  Stainless steel (and low-carbon steel) sheet-metal work compatible with vacuum requirements
  Vacuum and cryogenics technology
  Leak-tight welding techniques
  Leak detection with helium mass spectrometry
  …
The cryostat design engineer is confronted with a multidisciplinary environment in which he needs to master “a little of everything”
…not to forget the industrialisation aspects when he is asked to produce cryostats in large series

## Slide 134: Thank you for your attention!
Acknowledgements:
The work presented in this course is essentially the result of contributions from a number of colleagues and the work done during the design and construction of the LHC
I wish to acknowledge in particular for the material provided and for their contributions in preparing this course:
R.Bonomi, P.Cruikshank, Ph.Lebrun, Y.Leclercq, D.Ramos, A.Vande Crean and G.Vandoni

## Slide 135: References and selected bibliography
A.Bejan, Heat Transfer, J.Wiley & Sons, Inc
CRYOGENIE, SES APPLICATIONS EN SUPRACONDUCTIVITE, IIF/IIR 1995, Techniques de l’ingenieur.
Superconducting Magnets, M.Wilson, Oxford Science Publications
R.R.Conte, Éléments de Cryogénie, Masson & Cie, Éditeurs.
Steven W.Van Sciver, Helium Cryogenics, The International Cryogenics Monograph Series, Plenum Press.
K. Mendelssohn, The quest for absolute zero, McGraw Hill (1966)
R.B. Scott, Cryogenic engineering, Van Nostrand, Princeton (1959)
G.G. Haselden, Cryogenic fundamentals, Academic Press, London (1971)
R.A. Barron, Cryogenic systems, Oxford University Press, New York (1985)
B.A. Hands, Cryogenic engineering, Academic Press, London (1986)
S.W. van Sciver, Helium cryogenics, Plenum Press, New York (1986)
K.D. Timmerhaus & T.M. Flynn, Cryogenic process engineering, Plenum Press, New York (1989)
Proceedings of CAS School on Superconductivity and Cryogenics for Particle Accelerators and Detectors, Erice (2002)
  U. Wagner, Refrigeration
  G. Vandoni, Heat transfer
  Ph. Lebrun, Design of a cryostat for superconducting accelerator magnet
Proceedings of ICEC and CEC/ICMC conferences