## Slide 1
Introduction to Beam Instrumentation
CAS 2013
Trondheim, Norway
18th – 29th August, 2013
Dr. Rhodri Jones
Head of the CERN Beam Instrumentation Group

## Slide 2: Introduction
What do we mean by beam instrumentation?
  The “eyes” of the machine operators
    i.e. the instruments that observe beam behaviour
    An accelerator can never be better than the instruments measuring its performance!
What does work in beam instrumentation entail?
  Design, construction & operation of instruments to observe particle beams
  R&D to find new or improve existing techniques to fulfill new requirements
  A combination of the following disciplines
    Applied & Accelerator Physics; Mechanical, Electronic & Software Engineering
  A fascinating field of work!
What beam parameters do we measure?
  Beam Position
    Horizontal and vertical throughout the accelerator
  Beam Intensity (& lifetime measurement for a storage ring/collider)
    Bunch-by-bunch charge and total circulating current
  Beam Loss
    Especially important for superconducting machines
  Beam profiles
    Transverse and longitudinal distribution
  Collision rate / Luminosity (for colliders)
    Measure of how well the beams are overlapped at the collision point

## Slide 3: More Measurements
Machine Chromaticity
Machine Tune
QF
QF
QF
QD
QD
SF
SF
SF
SD
SD
Spread in the Machine Tune due to Particle Energy Spread
Controlled by Sextupole magnets
Characteristic Frequency
of the Magnetic Lattice
Given by the strength of the
Quadrupole magnets
Optics Analogy:
Achromatic incident light
[Spread in particle energy]
Lens
[Quadrupole]
Focal length is
energy dependent

## Slide 4: The Typical Instruments
Beam Position
  electrostatic or electromagnetic pick-ups and related electronics
Beam Intensity
  beam current transformers
Beam Profile
  secondary emission grids and screens
  wire scanners
  synchrotron light monitors
  ionisation and luminescence monitors
  femtosecond diagnostics for ultra short bunches
Beam Loss
  ionisation chambers or pin diodes
Machine Tune and Chromaticity
  in diagnostics section of tomorrow
Luminosity
  in diagnostics section of tomorrow

## Slide 5: Measuring Beam Position – The Principle
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-

## Slide 6: Wall Current Monitor – The Principle
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
V
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Ceramic Insert
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+

## Slide 7: Wall Current Monitor – Beam Response
L
IB
V
R
C
Frequency
Response
0
0
IB

## Slide 8: Electrostatic Monitor – The Principle
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-

## Slide 9: Electrostatic Monitor – Beam Response
VB
V
R
C
Frequency (Hz)
Response (V)
0
0
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+

## Slide 10: Electrostatic Monitor – The Principle
-
-
-
-
-
+
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
+
-
-
-
-
-
-
-
+
+
+
+
-
+
+
+
-
-
+
-
+
+
-
-
+
-
-
+
+
-
+
-
V
-
-
-
-
-
-
-
-
+
+
+
+
+
-
-
+
-
+
-
+
-
+
-
-
-
-
+
+
+
-
+
+
-
+
-
-
+
-
+
-
-
-
-
-
+
+
+
-
+
+
-
-
+
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+

## Slide 11: Electrostatic Pick-up – Button
Low cost  most popular
Non-linear
  requires correction algorithm
  when beam is off-centre
For Button with Capacitance Ce & Characteristic Impedance R0
Transfer Impedance:
Lower Corner Frequency:
Area A
r

## Slide 12: A Real Example – The LHC Button
Frequency (Hz)
Response (V)
0
0

## Slide 13: Improving the Precision for Next Generation Accelerators
Standard BPMs give intensity signals which need to be subtracted to obtain a difference which is then proportional to position
  Difficult to do electronically without some of the intensity information leaking through
    When looking for small differences this leakage can dominate the measurement
    Typically 40-80dB (100 to 10000 in V) rejection  tens micron resolution for typical apertures
Solution – cavity BPMs allowing sub micron resolution
  Design the detector to collect only the difference signal
    Dipole Mode TM11 proportional to position & shifted in frequency with respect to monopole mode
f / GHz
U / V
Frequency Domain
TM01
TM11
TM02
U~Q
U~Qr
U~Q
Courtesy of D. Lipka,
DESY, Hamburg
TM01
TM11
TM02

## Slide 14: Today’s State of the Art BPMs
Obtain signal using waveguides that only couple to dipole mode
  Further suppression of monopole mode
Prototype BPM for ILC Final Focus
  Required resolution of 2nm (yes nano!) in a 6×12mm diameter beam pipe
  Achieved World Record (so far!) resolution of 8.7nm at ATF2 (KEK, Japan)
Monopole Mode
Dipole Mode
Courtesy of D. Lipka,
DESY, Hamburg
Courtesy of D. Lipka & Y. Honda

## Slide 15: Criteria for Electronics Choice - so called “Processor Electronics”
Accuracy
  mechanical and electromagnetic errors
  electronic components
Resolution
Stability over time
Sensitivity and Dynamic Range
Acquisition Time
  measurement time
  repetition time
Linearity
  aperture & intensity
Radiation tolerance

## Slide 16: Processing System Families
D I G I T I Z E R
turn by turn
no turn by turn
Legend:
/ Single channel
Wide Band
Narrow band
Normaliser
Processor
Active
Circuitry
Heterodyne
POS = (A-B)
Synchronous
Detection
Automatic
Gain Control
on S
Multiplexed
Passive
Normalisation
POS = [log(A/B)]
= [log(A)-log(B)]
Differential
Amplifier
Logarithmic
Amplifiers
Individual
Treatment
Limiter,
Dt to Ampl.
Amplitude
to Time
POS = [A/B]
POS = [ATN(A/B)]
Amplitude
to Phase
.
Limiter,
f to Ampl.
POS = D / S
Heterodyne
Hybrid
D / S
Homodyne
Detection
Electrodes
A, B
Direct
Digitisation
POS = D / S

## Slide 17: Linearity Comparison

## Slide 18: Amplitude to Time Normalisation
A
B
1.5ns
B + 1.5ns
A
B
Beam
Splitter
Delay lines
Combiner
Pick-up

## Slide 19: Amplitude to Time Normalisation
A
B
A + (B + 1.5ns)
A
B
B + (A + 1.5ns)
Dt depends on position

## Slide 20: BPM Acquisition Electronics Amplitude to Time Normaliser
Advantages
Fast normalisation (< 25ns)
  bunch to bunch measurement
Signal dynamic independent of the number of bunches
  Input dynamic range ~45 dB
  No need for gain selection
Reduced number of channels
  normalisation at the front-end
~10 dB compression of the position dynamic due to the recombination of signals
Independent of external timing
Time encoding allows fibre optic transmission to be used
Limitations
Currently reserved for beams with empty RF buckets between bunches e.g.
  LHC 400MHz RF but 25ns spacing
  1 bunch every 10 buckets filled
Tight time adjustment required
No Intensity information
Propagation delay stability and switching time uncertainty are the limiting performance factors

## Slide 21: What one can do with such a System
Used in the CERN-SPS for electron cloud & instability studies

## Slide 22: The Typical Instruments
Beam Position
  electrostatic or electromagnetic pick-ups and related electronics
Beam Intensity
  beam current transformers
Beam Profile
  secondary emission grids and screens
  wire scanners
  synchrotron light monitors
  ionisation and luminescence monitors
  Femtosecond diagnostics for ultra short bunches
Beam Loss
  ionisation chambers or pin diodes
Machine Tunes and Chromacitities
  in diagnostics section of tomorrow
Luminosity
  in diagnostics section of tomorrow

## Slide 23: Current Transformers
Beam current
Magnetic field
ri
ro
Fields are very low
Capture magnetic field lines with cores of high
relative permeability
(CoFe based amorphous alloy Vitrovac: μr= 105)
w
N Turn winding
Transformer Inductance

## Slide 24: The Active AC transformer
CS
R
Beam signal
Transformer output signal
Winding of N turns and Inductance L
RF
RL
A
t
t
IB
U
L

## Slide 25: Fast Beam Current Transformer
500MHz Bandwidth
Low droop (< 0.2%/ms)
BEAM
Image
Current
Ceramic Gap
80nm Ti Coating
20W to improve
impedance
1:40 Passive
Transformer
Calibration winding

## Slide 26: Acquisition Electronics
FBCT Signal after 200m of Cable
Integrator Output
Data taken on LHC type beams at the CERN-SPS
25ns

## Slide 27: What one can do with such a System
Bad RF Capture of a single LHC Batch in the SPS (72 bunches)

## Slide 28: The DC transformer
B
I
AC transformers can be extended to very low frequency but not to DC ( no dI/dt ! )
DC measurement is required in storage rings
To do this:
  Take advantage of non-linear magnetisation curve
  Use 2 identical cores modulated with opposite polarities

## Slide 29: DCCT Principle – Case 1: no beam
I
B
Modulation Current - Core 1
Modulation Current - Core 2
IM
t
Hysteresis loop
of modulator cores

## Slide 30: DCCT Principle – Case 1: no beam
I
B
V
t
dB/dt - Core 1 (V1)
dB/dt - Core 2 (V2)
Output voltage = V1 – V2

## Slide 31: DCCT Principle – Case 2: with beam
Beam Current IB
V
t
IB
Output signal is at TWICE
the modulation frequency
dB/dt - Core 1 (V1)
dB/dt - Core 2 (V2)
Output voltage = V1 – V2
I
B

## Slide 32: Zero Flux DCCT Schematic
Beam
Compensation current Ifeedback = - Ibeam
Modulator
V = R  Ibeam
Power supply
R
Synchronous
detector
Va - Vb
Vb
Va

## Slide 33: The Typical Instruments
Beam Position
  electrostatic or electromagnetic pick-ups and related electronics
Beam Intensity
  beam current transformers
Beam Profile
  secondary emission grids and screens
  wire scanners
  synchrotron light monitors
  ionisation and luminescence monitors
  femtosecond diagnostics for ultra short bunches
Beam Loss
  ionisation chambers or pin diodes
Machine Tunes and Chromacitities
  in diagnostics section of tomorrow
Luminosity
  in diagnostics section of tomorrow

## Slide 34: Secondary Emission (SEM) Grids
When the beam passes through secondary electrons are ejected from the wires
The liberated electrons are removed using a polarisation voltage
The current flowing back onto the wires is measured
One amplifier/ADC chain is used for each wire

## Slide 35: Profiles from SEM grids
Charge density measured from each wire gives a projection of the beam profile in either horizontal or vertical plane
Resolution is given by distance between wires
Used only in low energy linacs and transfer lines as heating is too great for circulating beams

## Slide 36: Wire Scanners
A thin wire is moved across the beam
  has to move fast to avoid excessive heating of the wire and/or beam loss
Detection
  Secondary particle shower detected outside the vacuum chamber using a scintillator/photo-multiplier assembly
  Secondary emission current detected as for SEM grids
Correlating wire position with detected signal gives the beam profile
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+

## Slide 37: Beam Profile Monitoring using Screens
OTR Screen
Mirror
Intensifier -
CCD
Beam
Lens
Exit window
Optical Transition Radiation
  Radiation emitted when a charged particle beam goes through the interface of 2 media with different dielectric constants
  surface phenomenon allows the use of very thin screens (~10mm)

## Slide 38: Beam Profile Monitoring using Screens
Screen Types
  Luminescence Screens
    destructive (thick) but work during setting-up with low intensities
  Optical Transition Radiation (OTR) screens
    much less destructive (thin) but require higher intensity

## Slide 39: Beam Profile Monitoring using Screens
Usual configuration
  Combine several screens in one housing e.g.
    Al2O3 luminescent screen for setting-up with low intensity
    Thin (~10um) Ti OTR screen for high intensity measurements
    Carbon OTR screen for very high intensity operation
Advantages compared to SEM grids
  allows analogue camera or CCD acquisition
  gives two dimensional information
  high resolution: ~ 400 x 300 = 120’000 pixels for a standard CCD
  more economical
    Simpler mechanics & readout electronics
  time resolution depends on choice of image capture device
    From CCD in video mode at 50Hz to Streak camera in the GHz range

## Slide 40: Luminescence Profile Monitor
Beam
N2 ground state
e-
N2 excited state
Photon
emitted
Ion beam
Blackened walls
Vacuum gauge
Valve
Viewport
150mm flange
Lens, Image-Intensifier
and CCD Camera
N2-fluorescent gas
equally distributed

## Slide 41: Luminescence Profile Monitor
CERN-SPS Measurements
Profile Collected every 20ms
Local Pressure at ~510-7 Torr
2D
Side view
3D
Image
Beam Size
Time
Beam Size
Time
Injection
Beam size shrinks as
beam is accelerated
Fast extraction
Slow extraction

## Slide 42: The Synchrotron Light Monitor
Beam
Synchrotron Light from
Bending Magnet
or Undulator

## Slide 43: The Synchrotron Light Monitor

## Slide 44: Measuring Ultra Short Bunches
Next Generation FELs & Linear Colliders
  Use ultra short bunches to increase brightness or improve luminosity
How do we measure such short bunches?
  Transverse deflecting cavity
p+ @ LHC | 250ps
H- @ SNS | 100ps
e- @ ILC | 500fs
e- @ CLIC | 130fs
e- @ XFEL | 80fs
e- @ LCLS | <75fs
Destructive Measurement

## Slide 45: Electro-Optic Sampling – Non Destructive
Spectral Decoding
Temporal decoding
Limited to >250fs by laser bandwidth
Limited to >30fs by sampling laser pulse

## Slide 46: The Typical Instruments
Beam Position
  electrostatic or electromagnetic pick-ups and related electronics
Beam Intensity
  beam current transformers
Beam Profile
  secondary emission grids and screens
  wire scanners
  synchrotron light monitors
  ionisation and luminescence monitors
  femtosecond diagnostics for ultra short bunches
Beam Loss
  ionisation chambers or pin diodes
Machine Tunes and Chromacitities
  in diagnostics section of tomorrow
Luminosity
  in diagnostics section of tomorrow

## Slide 47: Beam Loss Detectors
Role of a BLM system:
  Protect the machine from damage
  Dump the beam to avoid magnet quenches (for SC magnets)
  Diagnostic tool to improve the performance of the accelerator
Common types of monitor
  Long ionisation chamber (charge detection)
    Up to several km of gas filled hollow coaxial cables
    Position sensitivity achieved by comparing direct & reflected pulse
      e.g. SLAC – 8m position resolution (30ns) over 3.5km cable length
    Dynamic range of up to 104
  Fibre optic monitors
    Similar layout with electrical signals replaced by light produced through Cerenkov effect

## Slide 48: Beam Loss Detectors
Common types of monitor (cont)
  Short ionisation chamber (charge detection)
    Typically gas filled with many metallic electrodes and kV bias
    Speed limited by ion collection time - tens of microseconds
    Dynamic range of up to 108
iin(t)
iin(t) + Iref
LHC

## Slide 49: Beam Loss Detectors
Common types of monitor (cont)
  PIN photodiode (count detection)
    Detect MIP crossing photodiodes
    Count rate proportional to beam loss
    Speed limited by integration time
    Dynamic range of up to 109
HERA-p
Comparator

## Slide 50: Beam Loss Detectors – New Materials
Diamond Detectors
  Fast & sensitive
  Used in LHC to distinguish bunch by bunch losses
  Investigations now ongoing to see if they can work in cryogenic conditions
Courtesy of E. Griesmayer

## Slide 51: BLM Threshold Level Estimation

## Slide 52: Summary
This was an overview of the common types of instruments that can be found in most accelerators
  Only a small subset of those currently in use or being developed with many exotic instruments tailored for specific accelerator needs existing
Tomorrow you will see how to use these instruments to run and optimise accelerators
  Introduction to Accelerator Beam Diagnostics (H. Schmickler)
Join the afternoon course:
Beam Instrumentation & Diagnostics
  For an in-depth analysis of all these instruments and on their application in various accelerators